CAIET DE SARCINI

la concursul pentru achiziționarea bunurilor

prin metoda cererii ofertelor de prețuri

Lot 1

Potentiostat/galvanostat

| Nr. d/o | Cod CPV | Denumirea bunurilor solicitate | Unitatea de măsură | Cantitatea | Valoarea estimată (lei) |
| --- | --- | --- | --- | --- | --- |
| 1 | 38500000-0 | Potentiostat/Galvanostat | set | 1 | 600000 |
| TOTAL |  |  |  |  | 600000 |

Specificarea tehnică deplină solicitată

Potentiostat/galvanostat, with a compliance voltage of 30 V and a bandwidth of 1 MHz and computer with advanced software for electrochemical research.

- Compatibility with having AUTOLAB systems and used software NOVA

- Warranty – min. 36 months.

- Manuals and other documentation in English or Romanian.

- All software capacitating functionality have to be supplied together with hardware (included into final price).

- It is mandatory: deliver of equipment, install and launch on site, training of staff (included into offer).

- Available techniques: potentiostat, galvanostat, electrochemical impedance spectroscopy.

- Minimal set of working modes:

- LSV (Linear Sweep Voltammetry)

- CV (Cyclic Voltammetry)

- SWV (Square Wave Voltammetry)

- Chronovoltammetry

- Chronopotentiometry

- Electrochemical impedance spectroscopy

- Modular system (possibility to create multihousing software-potentiostat or another units, such as high current boosters or electrochemical impedance spectra recording unit)

- Possibility to upgrade potentiostat with internal Electrochemical channel with Electrochemical Quartz Crystal Microbalance internal module or pH and Temperature monitoring internal module.

- Possibility to work with 2,3 and 4 electrodes

Main instrument technical data:

| Potentiostat/galvanostat |  |
| --- | --- |
| PC interface | USB |
| DC-Potential range | At least ±10 V |
| Compliance voltage range | 30 V |
| Potential resolution | 0.3 µV |
| Potential accuracy | min. 0.2 % |
| Current ranges | from 10 nA to at least 1 A possibility to enhance up to 10 A in the potential range at least ±10 V |
| Maximum current | 2 A |
| Current resolution | min. 0.0003% of current range |
| Current accuracy | min 0.2% of current range |
| Input impedance | 1 TOhm |
| Maximum bandwidth | 1 MHz |
| Electrode connections | 2, 3 and 4 |
| Scan rate | At least 1 mV/s to 1V/s |
| Pulse time | No less than 1 ms |
| Frequency range | At least from 10 Hz to 1 MHz |
| Signal types | 1 sine, 5 sines, 15 sines |
| AC amplitude in potentiostatic mode | 0.2 mV to 0.35 mV rms |
| Requirements for data presentation | Nyquist, Bode, Admittance, Mott-Schottky |
| Minimal software capabilities | Fit, simulation, Kramers-Kronig procedure |
| Cables | all necessary cables (power and interface) |
| Module for electrochemical impedance spectroscopy |  |
| Amplitude range in galvanostatic mode | 0.0002 to 0.35, steps of 0.000025 times applied current range |
| Amplitude range in potentiostatic mode | 0.2 mVrms to 0.35 Vrms with 0.025 mVrms step |
| Frequency range | 10 µHz - 32 MHz |
| Maximum frequency in combination with Autolab | 1 MHz |
| Frequency resolution | 0.003% |
| Input range | 10 V |
| Titration vessel with thermostat jacket |  |
| Height | 82 mm |
| Material | Borosilicate glass 3.3 |
| Outer diameter | 78 mm |
| Volume | 50 – 150 ml |
| Titration vessel lid | 5 openings |
| Support stand with support rod | Complete, with stand rod, total length 42 cm |
| Ag/AgCl reference electrode |  |
| Diaphragm | Ceramic diaphragm |
| Electrode plug-in head | Metrohm plug-in head B |
| Electrolyte outflow (µL/h) with reference to 3 mol/L KCl | 5...15 µL/h |
| Internal reference electrode type | LL system |
| Max. installation length | 125 mm |
| Min. immersion depth | 10 mm |
| Reference electrolyte type | c(KCl) = 3 mol/L |
| Reference resistance (kΩ) with reference to 3 mol/L KCl | <3 kΩ |
| Reference system | LL system |
| SGJ sleeve | Standard ground-joint 14/15 |
| Shaft diameter bottom | 5 mm |
| Shaft diameter top | 12 mm |
| Shaft material | Glass |
| Temperature range, long-term | 0 ... 80°C |
| Temperature range, short-term | 0 ... 80°C |
| Electrode cable | 2 x plug B (4 mm) for connecting electrodes with plug-in head B (length 1 m) |
| Computer (brand name) with software for electrochemical measurements |  |
| Platform | Minitower/Tower Brand Name, Producator deține certificat ISO9001 |
| CPU | Intel Core i5 min. Generatia 8; min 6 Core, min 6 Threads, frecvența de bază min. 2.8 Ghz, max. 14nm, cache min. 9Mb, TDP: max. 65W, integrated video |
| MB | Sokets compatibil cu CPU, RAM (2400Mhz) min 2 slots with support up to 32 GB, 1xDisplay port, 1xVGA, DVI, D-SUB sau HDMI port, PCI Express x16, M.2 PCIex1, 2xUSB 3.1 (front), 4xUSB 2.0 sau 3.0 (rear), 2xSATA 6.0Gb/s, 1xGbit LAN. |
| RAM | Min. 8 GB, DDR4-2400; 1 x 8 GB, frecvenţa min. 2400 MHz, compatibilă cu placa de bază |
| Video card | Integrated |
| SSD | min. 256 GB, M.2 PCIe NVMe |
| HDD | Min. 1 TB, 7200 rpm SATA III, Cache min. 32 MB. |
| Network | Integrated, Supports 10/100/1000Mbps |
| Sound card | Integrated |
| Ports | 2xUSB 3.1 (front), 4xUSB 2.0 sau 3.0 (rear) 1xDisplay port, 1xVGA, DVI, D-SUB sau HDMI port PS/2 optional, Audio, Lan RJ45 |
| Keyboard | PS/2 or USB, Eng/Rus |
| Mouse | PS/2 or USB, optical, scrolling |
| Speakers | External RMS 10W |
| Power Supply | min. 500W active PFC |
| Surge protector | 220V, 5 sockets, 5 m. |
| UPS | min. 650 VA, AVR, (2x Schuko sockets) |
| OS/Software | Microsoft Windows 10 Professional 64-bit OEM Software for electrochemical measurements |
| Monitor | Tip ecran – IPS LED sau echivalent. Formatul ecranului - Wide (formatul imaginii: 16:9). Mărime - Min. 27 inch; Rezoluţie - 1920 x 1080 pixeli; Timp de răspuns 5 ms; Reproducere culori - Min. 16 milioane; Raportul de contrast - Min. 1000:1 (static); 5000000:1 (dynamic) Luminozitate tipică - 250 ~ 300 cd/m; Unghi maxim de vizibilitate orizontală/verticală - 170°/170° USB porturi x2 Conectivitate – VGA, DisplayPort sau HDMI. Height Adjustment; Pivot; Accesorii - Livrat cu cablu de conexiune şi cablu de alimentare. |
| Warranty | Min. 36 months |