2. Radiologie intraorala - 1 buc

| Radiologie intraorala (caracteristici de baza) . |  |
| --- | --- |
| Descriere | Sistem de radiologie retroalveoiara cu un tub de ultima generaţie înalta tensiune 60/65/70 kV Curent anodic 4/8*mA Dimensiunea focala 0,4 mm Generator cu potenţial constant controlat de microprocesor. |

|  | Specificaţii |
| --- | --- |
| Generator | Constant potenţial, microprocessor-controlled |
| Frecventa de lucru | 145-230 KHz (typically 175 KHz) |
| Focal spot | 0,4 mm (iEC 336) |
| Curentul Anod | 4/8 mA |
| Voltage X-ray tube | 60 / 65 / 70 kV (*) |
| Timpul de expunere | . 0.020 - 1.000 seconds, R’10 and R’20 scale |
| Distanta de sursa | 20 e 30 cm |
| Distanta de iradiere | 35 x 45 mm rectangular, 0 55 mm and 0 60 mm round |
| Collimatoare adiţionale | 31x41 mm and 22 x 35 mm, for size 2 and size 1 sensors |
| Total filtrare | 2.0 mm Al @ 70kV |
| Power supply | 50/60 Hz, 230-240Vac ±10% |
| Duty Cycle | continuous operation with self-adjustment up to ls/80s total |

| Stability | Automatic lock/release, with touch-sensitive activation |
| --- | --- |
| Arms | Available in 3 lengths: 40 cm - 60cm - 90 cm |
| Max. arm extension | 230 cm, from wall /W |
| Certification | CE 0051, FDA approved |